UTX620   XRF Specification Sheet

1、Standard：DIN50987、ISO3497、ASTN.D568

2、Parameter：

| Testing Principle | EDXRF |
| --- | --- |
| Testing Objects | Solid/liquid/powder |
| Functions | Single layer testing、alloy layer testing、multi-layer testing、the limit is 5 layers； |
|  | Element Analysis |
|  | Coating liquid analysis(plating syrup main element analysis) |
| Testing Limit | Coating thickness：0.03 um，element Analysis：100ppm |
| Range of thickness | 0.03-35 um， |
| Error of thickness | 1st layer<±5%，2nd layer<±10%，3rd layer<±15% |
| Range of Element Analysis | Al13-U92， |
| Testing Time | Coating thickness：10-60s，element Analysis：10-60s |
| X Ray Relating Parts |  |
| X Ray tube | Made in USA，brand：OXFORD |
| X Ray tube target | W、Rh、Mo、Ag、Cr optional |
| X Ray tube window | Nylon |
| Electrical Parameters | 50W、4-50kV，0-1000uA |
| Way of cooling | Silicone Grease cooling |
| X ray Irradiation mode | ↓X ray from top down |
| Irradiation area | The smallest spotφ0.5mm |
| Lifetime | 15000-20000h |
| Detector |  |
| Detector | Proportional counter detector，brand：AMPTEK |
| Way of cooling | Electric Refrigeration (-1900C) |
| Resolution | Half-peak width of Mn 760eV |
| Detection window | Beryllium window size：400mm2 |
| High voltage power |  |
| Origin and brand | Made in USA，brand：SPELLMAN |
| Input voltage | 24VDC±10%、 |
| Output voltage | 0-50kV、0-2.0mA，Output accuracy 0.01% |
| Remarks | Have the function of automatic detection of software |
| Optical System |  |
| Collimator system | Single collimatorφ0.2mm、 |
| Focusing System | Laser Focus |
| Filter | Kinds of Filter (Ni、Cu、Mo、Al、Ti) |
| Sample observation system | 3,000,000-pixel high-definition CCD camera，20 optical zoom |
| Sample Room |  |
| Sample Room inner space（mm） | （W×H×D）：450×160×400mm |
| Dimension | （W×D）：250×275mm |
| Weight load of Samples table | 5kg |
| Dimension limit of sample | W×H×D：440×150×390mm |
| Controlling way of sample table | manual |
| Working table Moving technical data | X=200±5、Y=200±5、Z=160±5mm |
| Data dealing system |  |
| Controlling computer | Lenovo M7150，1G Memory，160G hard disk |
| Displayer | Lenovo 17 LCD displayer |
| printer | HP2468 color ink jet printer |
| result | Word、PDF |
| UPS power | 1 set |
| Software | Operating system：Windows XP、OEM Chinese Edition |
|  | Coating thickness testing software、element analysis software |
| Software language | Chinese\English\Others |
| Package parameter |  |
| Dimension （mm） | W×H×D：W 580×H 600×D 715mm |
| Weight | About 60kg |
| Attachment |  |
| CuAg standard calibration 1 pcs 、Ag 1 pcs |  |
| Working environment |  |
| Working environment | Tem 15-250C，humidity40-70%RH |
| Power system | Single 220V±10%，Working voltage in permitted range |
| Others | - Not near to devices that have strong magnetic field ,electric field, high-frequency - Reduce vibration - Less dust, low humidity, non-corrosive gas - No direct sunlight - As a response to the earthquake, it is necessary to consider the installation is fixed enough. |